## SPM Part 1
DENTAL INTRAORAL SAFETY AND PREVENTIVE MAINTENANCE INSPECTION REPORT
(Part 1)
EQUIPMENT IDENTIFICATION | [6 empty] | Date of Inspection:
Dental Clinic: | [8 empty] | Room No.:
Address: | [9 empty] | City:
Control Manufacturer: |  |  |  |  | Model No.: | [5 empty] | Serial No.:
Tube Manufacturer: |  |  |  |  | Model No.: | [5 empty] | Serial No.:
RADIOGRAPHIC INSPECTION REPORT
A. X-RAY TUBE AND SUPPORT SYSTEM | [5 empty] | RESULT |  | REMARKS
Cones/Collimator/Filters
Angulation Indicator
Indicator Lights
Security of Mounting (overall)
Stability of Motion (drift)
Electrical Cables/Connections/Grounds
B. CONTROL/TRANSFORMER
Indicator Lights & Meters
Electrical Cables/Connections/Grounds
Overload Protection Circuits
Control Knobs & Buttons
Audible Exposure Indicator
Exposure Control Switch
Copy of Report Sent to Dental Office?
By submitting this report as a qualified person, I ensure that this ionizing equipment and apparatus is in safe operating
condition, has undergone a radiation calibration, and the results of all testing fall within acceptable tolerances.
|  | Version 1.3 |  | SERVICE ENGINEER:
|  | 2021-01-01 |  | Saskatchewan Journeyman Licence No.:
CLEAR FORM
SAVE SPM FORM

## SPM Part 2
DENTAL SAFETY AND PREVENTIVE MAINTENANCE INSPECTION REPORT
(Part 2) | [25 empty] | PASS
FAIL
EQUIPMENT IDENTIFICATION | [13 empty] | DATE OF INSPECTION: | [11 empty] | N/A
DENTAL CLINIC: | [15 empty] | Room No.:
RADIOGRAPHIC CALIBRATION REPORT
A. PATIENT ENTRANCE EXPOSURE
| Take one Exposure Reading at the cone exit using clinic Adult Bitewing setting. | [22 empty] | Result:
kV: |  | mA: |  |  |  | Time: |  | (s) | Source-to-Cone Tip Distance: | [9 empty] | (in)
Monitoring Instrument(s): | [11 empty] | Exposure: |  |  |  | (mR) |  | Technology:
B. TIMER CHECK: | [5 empty] | For time < 20 p or 0.33 sec, Difference ≤ 17 ms or 1 pulse OTHERWISE Difference ≤ 5%
Set Time |  |  |  |  | Calc Set |  |  | Measured Time | [6 empty] | Calc Meas |  |  | Difference
(s) OR |  | (Pulse) |  |  | Time (s) |  |  | (s) OR |  |  |  | (Pulse) |  |  | Time (s) |  |  | (ms) OR |  |  | (%)
Result:
Comments:
C. LINEARITY CHECK: | [15 empty] | Cone Tip-to-Detector Distance: | [5 empty] | (cm)
Set mA |  | Set Time |  |  |  |  | Calc Set |  |  | Measured |  |  |  | Measured |  | mR/mAs
|  | (s) OR |  | (Pulse) |  |  | Time (s) |  |  | kVp |  |  |  | mR
Result:
[5 empty] | Highest - Lowest |  |  |  |  | = |  | 0 |  |  |  | (must be < 0.10). Requires 3 or more readings to create a pass.
Highest + Lowest
Comments:
D. kVp CHECK
Set kVp |  |  | Measured kVp |  |  |  | Set Time |  |  |  | % kV Difference
Result:
kVp difference < 7%
Comments:
E. HALF VALUE LAYER DETERMINATION: | [12 empty] | kVp:
|  |  | mm Al | mR |  |  | Log10(mR) | [5 empty] | 57 kVp: HVL greater than 1.5 mm Al
|  | 0 | [10 empty] | 60 or 70 kVp: HVL greater than 2.0 mm Al | [10 empty] | Result:
The kVp field is required to assess a "Pass" or "Fail"
[13 empty] | HVL = |  | 0 |  | mm Al
Comments:
F. ADJUSTMENTS MADE:
WARNING: All tests must either have a pass or a comment indicating why the test failed prior to submitting. Forms without a pass or comment will not be approved.
RSU Equipment No.:
(to be completed by RSU office)
[20 empty] | Version 1.3 |  |  | 2021-01-01
: //
CLEAR FORM
SAVE SPM FORM
Access Entry

## README
DENTAL SAFETY AND PREVENTIVE MAINTENANCE FORM
Welcome to the RSU dental SPM form. The purpose of the SPM testing is to inform the owner that their X-ray equipment is in safe operating condition. The testing is based on the standards of Health Canada Safety Code 30. Equipment that cannot meet these standards may compromise image quality and/or patient dose. Instructions for each test are provided in the purple tabs of this form. SPM inspections are required under sections 13 and 14 of The Radiation Health and Safety Regulations, 2005. The parts of the legislation that are releveant to dental X-Ray units are outlined below:
Safety and preventive maintenance inspections
13(1) An owner of ionizing radiation equipment and associated apparatus that is
used in a hospital or medical clinic must arrange for the inspection of that
equipment and apparatus by a qualified person in a manner and to a degree that is
satisfactory to an officer to ensure that the equipment and apparatus:
(a) is in safe operating condition; and
(b) has undergone a radiation calibration, the results of which are recorded on a
form supplied by the department.
(2) For the purposes of subsection (1), a person who holds a valid restricted x-ray
journeyperson’s licence issued pursuant to The Electrical Licensing Act is a
qualified person.
(3) An owner of ionizing radiation equipment that is used in a dental, chiropractic
or veterinary clinic must arrange for the inspection of the equipment in accordance
with subsection (1).
(4) A person who conducts an inspection pursuant to subsection (1) or (3) shall,
within 30 days after completing the inspection, submit to the Radiation Safety Unit
on the form mentioned in clause (1)(b) details of all tests carried out and all
measurements made in the course of the inspection.
Frequency of inspections
14(5) Subject to subsections (6) and (7), an inspection required by subsection 13(3) is
to be carried out not less than:
(a) once every three years for dental or chiropractic x-ray equipment; and
(b) once every five years for veterinary x-ray equipment.
(6) No inspection is required pursuant to subsection 13(3) until five years have
elapsed since the date of manufacture of the equipment.
(7) In the case of chiropractic x-ray equipment 15 years of age or older, an
inspection required by subsection 13(3) is to be carried out not less than once per
year.
Complete the form as follows: 1) Fill out the form as completely as possible. - Only the blue fields are required. Calculations are built in. - For each test, either a "PASS" or comment on "FAIL" is required. - Incomplete SPMs will be rejected and returned to the owner. 2) Once complete, hit the SAVE FORM button at the bottom of either page or click "Save As" in the "File" menu. - The save will not complete unless the form is adequately filled out. 3) Submit the saved form via email to radiationsafety@gov.sk.ca. - Submission of the competed form is required within 30 days of inspection.
Track Change Summary
Version | Release Date
1 | 2016-01-01 | u First Release
u Replace old non-macro forms
u Added preamble message to user
u Locked all fields not used for testing
u Introduced Calculations and Pass/Fail Results
u Created function buttons that save and clear the form
u Created function button that enters SPM into Access database and saves file to ORS
u Introduced quality checks into Save button and hid ribbon to force users to use quality checks
1.1 | October, 2016 | u Minor fix in the HVL Pass/Fail criteria
u Re-introduced Ribbon and wrote code that forces quality check when user uses the Save As function in the file menue
u Fixed issue for users with 2007 version of Excel that would not compile. Issue with TRIM and FORMAT functions. For that compiler, the reference needs to be strings.Trim and strings.Format to work.
1.2 | June, 2017 | u Added README tab with Track Change Summary
u Removed the Welcome Message pop-up and placed it in the README tab to avoid users from seeing it every time.
u Removed code reference to Outlook Library to avoid compiler error created when users with future versions of Excel save the file
u Added Entrance Exposure Test to determine Patient Exposure
u Instruction Tabs added explaining each test procedure
1.3 | January, 2021 | u Updated Logo for Visual Identity
u Changed date format to MMMM dd, yyyy for better clarity

## Entrance Exposure Check
Patient Entrance Exposure Check Instructions | [5 empty] | Setup with a solid state detector with lead backing
Purpose
The purpose of the patient entrance exposure check is to ensure that the clinically used adult bitewing does not produce excessive patient exposures for the imaging technology being used.
Procedure
1. Consult with the clinical staff to aquire the adult bitewing technique (kVp, mA, time) for the X-Ray unit being tested.
2. Record the source-to-cone distance (typically 8" or 12") in the provided field on the form.
3. Enter the clinically used kVp, mA and time into the respective fields on the form.
For X-Ray units that display timer settings in seconds, record the set timer settings in the time field (in units of seconds).
For X-Ray units that display timer settings in pulses, convert pulses to seconds by dividing the number of pulses by 60 and record the set timer settings in the time field (in units of seconds).
4. Place your detector on a minimally reflective surface (i.e., treatment chair) to minimize backscatter.
5. Place the active area of your detector directly against the treatment head cone.
NOTE: DO NOT LEAVE ANY SPACE BETWEEN THE DETECTOR ACTIVE AREA AND THE CONE TIP.
Setup with an ion chamber
6. Take an exposure and record the result (in milliroentgen) in the exposure field.
Warning Criteria
Different imaging technologies (film, digital psp scanned plates, digital sensors etc…) require different amounts of radiation to produce an acceptable image. The following table indicates the exposure criteria that would trigger a warning of possible excessive patient entrance exposures.
Upper Exposure Warning Limit (mR)
| kVp: | 60 | 65 | 70
| FILM | 220 | 190 | 165 | (Note: The film Limits are for E and F-Speed film only. Sites using D-Speed film are strongly encouraged to switch to F-Speed film.)
| DIGITAL PSP PLATES | 220 | 190 | 165
| DIGITAL SENSORS | 195 | 165 | 140
Note: The limits in the above table are quite liberal. The majority of dental sites in Saskatchewan are able to achieve patient entrance exposures below these limits. Film limits were derived from Radiation Protection in Dentistry - Recommended Safety Procedures for the Use of Dental X-Ray Equipment - Safety Code 30. These same limits were adopted for digital PSP plates. Based on RSU data, digital sensors require lower amounts of radiation to produce a proper image so the Safety Code limits were reduced by a modest 25 mR.
Actions
If the warning result appears, there may be an issue with excessive patient exposures. To address this issue, the adult bitewing programmed timer setting could be reduced. Consider stepping down the timer setting and assessing the image quality at each step. Endeavour to find the lowest timer setting that does not compromise image quality.

## Timer Check
Timer Check Instructions |  | Setup with a solid state detector with lead backing
Purpose
The purpose of the timer check is to ensure that set exposure time accurately represent the exposure time produced by the X-Ray unit. This check is important to ensure that patient exposures are not higher or lower than intended by the operator. The intent is to prevent issues with image quality and patient dose (through high exposure or unneccessary repeats).
Procedure
1. Select a set of four timer settings covering the entire range of clinically used timer settings.
For X-Ray units that display timer settings in seconds, record the set timer settings in the first column (in units of seconds).
For X-Ray units that display timer settings in pulses, record the set timer settings in the second column (in units of pulses).
2. Place the active area of your detector directly under the treatment head cone.
3. Measure the exposure time for the four timer settings chosen.
For X-Ray units that display timer settings in seconds, record the measured exposure times in the fourth column (in units of seconds).
For X-Ray units that display timer settings in pulses, record the measured exposure times in the fifth column (in units of pulses).
Pass/Fail Criteria
Setup with an ion chamber
For timer settings £ 20 pulses or 0.33 seconds, the set and measured times should not differ by more than 1 pulse or 0.017 seconds.
For timer settings > 20 pulses or 0.33 seconds, the set and measured times should not differ by more than 5%.
Actions
If the test results is a fail, make every effort to adjust the X-Ray unit to produce a passing result and repeat the test.
If the unit cannot be adjusted to produce a pass, please summarize what actions were taken in the comments section.

## Linearity Check
Linearity Check |  | Setup with a solid state detector with lead backing
Purpose
The purpose of the linearity check is to ensure that the exposure produced by the X-Ray unit properly scales with timer setting (e.g., a doubling of timer setting produces double the generated X-Rays) and/or tube current. A failure in the linearity check can be an indicator that an increased load on the X-Ray tube is causing a diminished output (tube failure).
Procedure
1. Select the clinically used tube current and enter it in the first column. (If more than one tube current is used clinically, test both settings).
2. Select a set of four timer settings covering the entire range of clinically used timer settings.
For X-Ray units that display timer settings in seconds, record the set timer settings in the second column (in units of seconds).
For X-Ray units that display timer settings in pulses, record the set timer settings in the third column (in units of pulses).
3. Place the active area of your detector directly under the treatment head cone.
4. Measure and record the tube voltage (column 5) and X-Ray exposure (column 6 in milliroentgen) for the four timer settings chosen.
Pass/Fail Criteria |  |  | Setup with an ion chamber
The calculated variance in linearity should not exceed 0.10 (10%).
Actions
If the test results is a fail, make every effort to adjust the X-Ray unit to produce a passing result and repeat the test.
If the unit cannot be adjusted to produce a pass, please summarize what actions were taken in the comments section.

## kVp Check
kVp Check |  | Setup with a solid state detector with lead backing
Purpose
The purpose of the kVp check is to ensure that the X-Ray unit is capable of repeatedly producing the peak voltage(s) it is designed to produce. Peak tube voltages of 60 or 70 kVp are typically chosen to produce the best anatomic contrast in dental X-Ray images. A failure in the kVp check can be an indicator of comprimised image quality.
Procedure
1. Enter the peak tube voltage clinically used in the first column. (If more than one tube current is used clinically, test both settings).
2. Select a set of three timer settings covering the entire range of clinically used timer settings.
For X-Ray units that display timer settings in seconds, record the set timer settings in the third column (in units of seconds).
For X-Ray units that display timer settings in pulses, record the set timer settings in the third column (in units of pulses).
3. Place the active area of your detector directly under the treatment head cone.
4. Measure and record the tube voltage (column 2) for the three timer settings chosen.
Pass/Fail Criteria |  |  | Setup with an ion chamber
The measured peak voltage should be within 7% of the set voltage.
Actions
If the test results is a fail, make every effort to adjust the X-Ray unit to produce a passing result and repeat the test.
If the unit cannot be adjusted to produce a pass, please summarize what actions were taken in the comments section.

## HVL Check
Beam Quality (HVL) Instructions
Purpose
The purpose of the HVL check is to ensure that the beam quality or penetrating power of the X-Ray beam is sufficient to avoid image quality issues or excessive patient skin doses.
Procedure
1. Place the active area of your detector at a fixed distance from the treatment cone.
Avoid placing non lead backed detectors on strong backscatter surfaces. Suspend detector if necessary.
2. Set the kVp that is clinically used.
3. Record the kVp value on the form in the field provided.
4. Take an exposure with no attenuation added and record the exposure value in the table provided.
5. Choose two thicknesses of aluminum (ideally, one thickness below the anticipated HVL and one thickness above).
6. Insert each thickness between the detector and the treatment cone head and take an exposure.
7. Record the exposure values in the table provided.
Pass/Fail Criteria
The HVL should be above 2.0 mm Al for 60kVp to 70 kVp units.
The HVL should be above 1.5 mm Al for 57 kVp units.
Actions
If the test results is a fail, make every effort to adjust the X-Ray unit to produce a passing result and repeat the test.
If the unit cannot be adjusted to produce a pass, please summarize what actions were taken in the comments section.
Notes
A kVp value must be entered in order for a pass/fail to appear.
Although not prohibited, the use of a survey meter to directly measure HVL is discouraged as this method has been found to be less accurate compared to measurements with aluminum attenuation as described in this procedure. If a survey meter is used to directly measure the HVL, enter the measured value in the comment field provided. You may then ignore the Pass/Fail warning when submitting the results.
Aluminum Plates
Detector or Chamber
Chair or Countertop
Treatment Head