## SPM Part 3
MEDICAL SAFETY AND PREVENTIVE MAINTENANCE INSPECTION REPORT
(Part 3) Fluoroscopy
EQUIPMENT IDENTIFICATION | [8 empty] | DATE OF INSPECTION:
FACILITY NAME: | [12 empty] | Room No.:
ADDRESS: | [10 empty] | CITY: |  |  |  |  | PC:
Image Intensifier Maker: | [6 empty] | Model: | [6 empty] | Serial No.:
Fluoro Tube Maker: | [6 empty] | Model: | [6 empty] | Serial No.:
FLUOROSCOPY INSPECTION REPORT
A. X-RAY TUBE AND SUPPORT SYSTEM | [5 empty] | RESULT |  |  |  |  | REMARKS
High Tension Cables
Shutter, auto collimation PBL
Support Cables
Locks
Alignment to Image Intensifier
Fluoro Timer
Foot Switch
Indicator Lights
Security of Mounting (overall) Electrical
Cables/Connections/Grounds
B. Image Intensifier and Tower
Indicator Lights & Meters
Electrical Cables/Connections/Grounds
Power Assists & Locks
Control Switch and Buttons
Compression Device
Automatic Brightness Control
Parking Mechanism
Support Cables
Lead Drapes
C. T.V. Monitor
Electrical Cables/Connections/Grounds
Control Switches and Buttons
Suspension System and Cables
D. Spot Film Device and Camera
Locks & Cables
Indicator Lights
Electrical Cables/Connections/Grounds
Microswitches and Programming
Grid Condition
E. Radiation Dose Rate and Resolution |  |  |  |  | Monitoring Instrument:
Maximum Radiation Dose Rate: | [6 empty] | R/minute | [7 empty] | kVp: |  | mA:
Resolution - Lucite QA Phantom: |  |  |  | - |  | - |  | - |  | - |  | No. of Holes |  |  | kVp: |  | mA:
Radiation Dose Rate Lucite Phantom: | [6 empty] | R/minute | [7 empty] | kVp: |  | mA:
SERVICE ENGINEER:
Saskatchewan Journeyman Licence No.:
RSU Equipment No.:
| Version 1.3 |  | (to be completed by RSU office)
2021-01-01
CLEAR FORM
SAVE FORM
ACCESS ENTRY

## README
FLUOROSCOPIC SAFETY AND PREVENTIVE MAINTENANCE FORM
Welcome to the RSU dental SPM form. The purpose of the SPM testing is to inform the owner that their X-ray equipment is in safe operating condition. The testing is based on the standards of Health Canada Safety Code 30. Equipment that cannot meet these standards may compromise image quality and/or patient dose. Instructions for each test are provided in the purple tabs of this form. SPM inspections are required under sections 13 and 14 of The Radiation Health and Safety Regulations, 2005. The parts of the legislation that are releveant to dental X-Ray units are outlined below:
Safety and preventive maintenance inspections
13(1) An owner of ionizing radiation equipment and associated apparatus that is
used in a hospital or medical clinic must arrange for the inspection of that
equipment and apparatus by a qualified person in a manner and to a degree that is
satisfactory to an officer to ensure that the equipment and apparatus:
(a) is in safe operating condition; and
(b) has undergone a radiation calibration, the results of which are recorded on a
form supplied by the department.
(2) For the purposes of subsection (1), a person who holds a valid restricted x-ray
journeyperson’s licence issued pursuant to The Electrical Licensing Act is a
qualified person.
(3) An owner of ionizing radiation equipment that is used in a dental, chiropractic
or veterinary clinic must arrange for the inspection of the equipment in accordance
with subsection (1).
(4) A person who conducts an inspection pursuant to subsection (1) or (3) shall,
within 30 days after completing the inspection, submit to the Radiation Safety Unit
on the form mentioned in clause (1)(b) details of all tests carried out and all
measurements made in the course of the inspection.
Frequency of inspections
14(5) Subject to subsections (6) and (7), an inspection required by subsection 13(3) is
to be carried out not less than:
(a) once every three years for dental or chiropractic x-ray equipment; and
(b) once every five years for veterinary x-ray equipment.
(6) No inspection is required pursuant to subsection 13(3) until five years have
elapsed since the date of manufacture of the equipment.
(7) In the case of chiropractic x-ray equipment 15 years of age or older, an
inspection required by subsection 13(3) is to be carried out not less than once per
year.
Complete the form as follows: 1) Fill out the form as completely as possible. - Only the blue fields are required. Calculations are built in. - For each test, either a "PASS" or comment on "FAIL" is required. - Incomplete SPMs will be rejected and returned to the owner. 2) Once complete, hit the SAVE FORM button at the bottom of either page or click "Save As" in the "File" menu. - The save will not complete unless the form is adequately filled out. 3) Submit the saved form via email to radiationsafety@gov.sk.ca. - Submission of the competed form is required within 30 days of inspection.
Track Change Summary
Version | Release Date
1 | 2016-01-01 | u First Release
u Replace old non-macro forms
u Added preamble message to user
u Locked all fields not used for testing
u Introduced Calculations and Pass/Fail Results
u Created function buttons that save and clear the form
u Created function button that enters SPM into Access database and saves file to ORS
u Introduced quality checks into Save button and hid ribbon to force users to use quality checks
1.1 | October, 2016 | u Re-introduced Ribbon and wrote code that forces quality check when user uses the Save As function in the file menue
u Fixed issue for users with 2007 version of Excel that would not compile. Issue with TRIM and FORMAT functions. For that compiler, the reference needs to be strings.Trim and strings.Format to work.
1.2 | June, 2018 | u Added README tab with Track Change Summary
u Removed the Welcome Message pop-up and placed it in the README tab to avoid users from seeing it every time.
u Removed code reference to Outlook Library to avoid compiler error created when users with future versions of Excel save the file
1.3 | January, 2021 | u Updated logo for visual identity
u Changed date format to mmmm dd, yyyy to avoid confusion
u Changed save filename to add an "_F" to distinguish Part 3 from Parts 1 and 2 submissions