## SPM Part 1
MEDICAL SAFETY AND PREVENTIVE MAINTENANCE INSPECTION REPORT
(Part 1)
EQUIPMENT IDENTIFICATION | [6 empty] | DATE OF INSPECTION:
FACILITY NAME: | [8 empty] | Room No.:
ADDRESS: | [7 empty] | CITY: |  |  | PC:
Control Manufacturer: | [5 empty] | Model: |  |  |  | Serial No.:
Tube Insert Manufacturer: | [5 empty] | Model: |  |  |  | Serial No.:
RADIOGRAPHIC INSPECTION REPORT
A. X-RAY TUBE AND SUPPORT SYSTEM |  |  |  |  | RESULT |  |  | REMARKS
High Tension Cables
SID Indicator
Support Cables
Locks
Light Field Alignment/PBL
Angulation Indicator
Bucky Centre Light
Indicator Lights
Security of Mounting (overall)
Ease of Motion (longitudinal, transverse, vertical)
Electrical Cables/Connections/Grounds
B. CONTROL/GENERATOR/TRANSFORMER
Indicator Lights & Meters
Electrical Cables/Connections/Grounds
Overload Protection Circuits
Control Knobs & Buttons
High Tension Cables
Exposure Control Switch
C. X-Ray Table
Anchor Bolts
Cables
Motor Drives and Limits
Bucky & Cassette Locks
Float & Power Top Switches
Auto Exposure Control Device
Bucky Slot Cover
Bucky Centred to Tube
Electrical Cables/Connections/Grounds
Grid Condition
Head Shoulder and Foot Rests
D. Wall Bucky
Locks & Cables
Grid Condition
Ease of Movement
Grid Centred to Tube
Electrical Cables/Connections/Grounds
Auto Exposure Control Device
Anchor Bolts/Mounting
E. Mobile Equipment
Drive Mechanism/Brakes
Battery Condition
Ease of Movement
SERVICE ENGINEER:
Saskatchewan Journeyman Licence No.:
| Version 1.3 | 2021-01-01
CLEAR FORM
SAVE FORM
CLEAR FORM
SAVE FORM

## SPM Part 2
MEDICAL SAFETY AND PREVENTIVE MAINTENANCE INSPECTION REPORT
(Part 2) | [24 empty] | PASS
FAIL
EQUIPMENT IDENTIFICATION | [12 empty] | DATE OF INSPECTION: | [11 empty] | N/A
FACILITY: | [16 empty] | Room No.:
Tube: same as part 1
RADIOGRAPHIC CALIBRATION REPORT
A. TIMER CHECK: |  |  | kVp: |  |  |  |  | mA: |  |  | Monitoring Instrument
Set Time (s) |  | Measured Time (s) | [5 empty] | Difference (%) |  |  |  | Set Time (s) |  |  |  |  | Measured Time (s) |  |  |  | Difference (%)
Result:
Comments: | [19 empty] | Criteria: Difference must be ≤ 5%
B. mA/mAs LINEARITY CHECK: | [5 empty] | kVp: |  |  |  |  | Monitoring Instrument
Set mA OR mAs |  | Set Time (s) |  |  |  |  | Measured kVp |  |  |  | Measured mR | [5 empty] | mR/mAs
Result:
[5 empty] | Highest - Lowest |  |  |  |  | = | 0 | [5 empty] | Criteria: (must be < 0.10). Requires 3 or more readings to create a pass.
Highest + Lowest
Comments:
C. kVp CHECK |  | (for all clinically used) | [8 empty] | Monitoring Instrument
Set kVp | Measured kVp |  |  |  | Set mA |  |  |  | % kV Difference |  |  | Set kVp |  |  |  | Measured kVp |  |  | Set mA |  |  | % kV Difference
Result:
Comments: | [13 empty] | Criteria: kVp difference < 7% for all clinically used settings.
D. HALF VALUE LAYER DETERMINATION: 80 kVp | [10 empty] | Monitoring Instrument
|  | mm Al |  | mR |  |  | Log10(mR)
| 0 | [22 empty] | Result:
[12 empty] | HVL = |  |  | 0 |  |  | mm Al
Comments: | [20 empty] | Criteria: > 2.9 mm Al
E. ADJUSTMENTS MADE:
WARNING: All tests must either have a pass or a comment indicating why the test failed prior to submitting. Forms without a pass or comment will not be approved.
RSU Equipment No.:
(to be completed by RSU office)
[21 empty] | Version 1.2 |  | 2017-06-01
CLEAR FORM
SAVE FORM
CLEAR FORM
SAVE FORM
Access Entry

## README
MEDICAL SAFETY AND PREVENTIVE MAINTENANCE FORM - GENERAL RADIOGRAPHY
Welcome to the RSU general radiography SPM form. The purpose of the SPM testing is to inform the owner that their X-ray equipment is in safe operating condition. The testing is based on the standards of Health Canada Safety Code 35. Equipment that cannot meet these standards may compromise image quality and/or patient dose. Complete the form as follows: 1) Fill out the form as completely as possible. - Only the blue fields are required. Calculations are built in. - For each test, either a "PASS" or comment on "FAIL" is required. - Incomplete SPMs will be rejected and returned to the owner. 2) Once complete, hit the "SAVE FORM" button at the bottom of either page or click "Save As" in the "File" menu. - The save will not complete unless the form is adequately filled out. 3) Submit the saved form via email to radiationsafety@gov.sk.ca. - Submission of the competed form is required within 30 days of inspection.
Track Change Summary
Version | Release Date
1 | 2016-01-01 | u First Release
u Replace old non-macro forms
u Added preamble message to user
u Locked all fields not used for testing
u Introduced Calculations and Pass/Fail Results
u Created function buttons that save and clear the form
u Created function button that enters SPM into Access database and saves file to ORS
u Introduced quality checks into Save button and hid ribbon to force users to use quality checks
1.1 | October, 2016 | u Minor fix in the HVL Pass/Fail criteria
u Re-introduced Ribbon and wrote code that forces quality check when user uses the Save As function in the file menue
u Fixed issue for users with 2007 version of Excel that would not compile. Issue with TRIM and FORMAT functions. For that compiler, the reference needs to be strings.Trim and strings.Format to work.
1.2 | 2017-06-01 | u Added README tab with Track Change Summary
u Removed the Welcome Message pop-up and placed it in the README tab to avoid users from seeing it every time.
u Removed code reference to Outlook Library to avoid compiler error created when users with future versions of Excel save the file
u Modify mA Linearity Test to be mA/mAs Linearity and changed mA(s) column to mA or mAs. Users need to enter unity in second column for mAs driven machines. Trying to clarify this.
u Instruction Tabs added explaining each test procedure
1.3 | 2021-01-01 | u Updated Logo on front page
u Changed date format to MMMM, DD, YYYY to avoid confusion

## Timer Check
Timer Check Instructions
Purpose
The purpose of the timer check is to ensure that set exposure time accurately represent the exposure time produced by the X-Ray unit. This check is important to ensure that patient exposures are not higher or lower than intended by the operator. The intent is to prevent issues with image quality and patient dose (through high exposure or unneccessary repeats).
Note
For mAs driven systems, this test cannot be performed. Please note that the system is mAs driven in the comment section on the test form.
Procedure
1. Select a set of up to twelve timer settings covering the entire range of clinically used timer settings.
2. Record the timer settings in the first column of the form.
3. Place a steel plate or lead apron on the table to protect the detector area (for digital systems only).
4. Place your detector directly under the X-Ray head and collimate the beam to the detector active area.
5. Measure the exposure time for the timer settings chosen.
6. Record the measured time in the second column of the form.
Pass/Fail Criteria
The set and measured times should not differ by more than 5%.
Actions
If the test results is a fail, make every effort to adjust the X-Ray unit to produce a passing result and repeat the test.
If the unit cannot be adjusted to produce a pass, please summarize what actions were taken in the comments section.

## Linearity Check
Linearity Check
Purpose
The purpose of the linearity check is to ensure that the exposure produced by the X-Ray unit properly scales with changes in mA, time or mAs (e.g., a doubling of mA, time or mAs setting produces double the generated X-Rays). A failure in the linearity check can be an indicator that an increased load on the X-Ray tube is causing a diminished output (tube failure).
Procedure
For mA and timer driven systems:
1. Select a clinically used range of tube current and timer setting combinations and enter them in the first two columns.
2. Select a single kVp setting used clinically and enter that value in the third column.
3. Place a steel plate or lead apron on the table to protect the detector area (for digital systems only).
4. Place your detector directly under the X-Ray head and collimate the beam to the detector active area.
5. Measure and record the X-Ray exposure in the fourth column for the settings chosen.
For mAs driven systems:
1. Select a clinically used range of mAs settings and enter them in the first column.
2. In the second column enter a value of "one" for each mAs setting. (The calculated mR/mAs result multiplies the first and second column values assuming an mA and time product. Since the first column is already an mA and time product, the second column is set to "1" to multiply that value by unity and preserve the mAs value).
2. Select a single kVp setting used clinically and enter that value in the third column.
3. Place a steel plate or lead apron on the table to protect the detector area (for digital systems only).
4. Place your detector directly under the X-Ray head and collimate the beam to the detector active area.
5. Measure and record the X-Ray exposure in the fourth column for the settings chosen.
Pass/Fail Criteria
The calculated variance in linearity should not exceed 0.10 (10%).
Actions
If the test results is a fail, make every effort to adjust the X-Ray unit to produce a passing result and repeat the test.
If the unit cannot be adjusted to produce a pass, please summarize what actions were taken in the comments section.

## kVp Check
kVp Check
Purpose
The purpose of the kVp check is to ensure that the X-Ray unit is capable of repeatedly producing the peak voltage(s) it is designed to produce. A failure in the kVp check can be an indicator of comprimised image quality.
Procedure
1. Enter the range of peak tube voltages clinically used in the first column.
2. Place a steel plate or lead apron on the table to protect the detector area (for digital systems only).
3. Place your detector directly under the X-Ray head and collimate the beam to the detector active area.
4. Measure and record the tube voltage (column 2) for each set kVp chosen.
5. Record the mA or mAs setting in the third column.
Pass/Fail Criteria
The measured peak voltage should be within 7% of the set voltage.
Actions
If the test results is a fail, make every effort to adjust the X-Ray unit to produce a passing result and repeat the test.
If the unit cannot be adjusted to produce a pass, please summarize what actions were taken in the comments section.

## HVL Check
Beam Quality (HVL) Instructions
Purpose
The purpose of the HVL check is to ensure that the beam quality or penetrating power of the X-Ray beam is sufficient to avoid image quality issues or excessive patient skin doses.
Procedure
1. Place a steel plate or lead apron on the table to protect the detector area (for digital systems only).
2. Set the tube voltage to 80 kVp (or the closest value if 80 is not available).
3. Place your detector directly under the X-Ray head and collimate the beam to the detector active area.
4. Take an exposure with no attenuation added and record the exposure value in the table provided.
5. Choose three thicknesses of aluminum (ideally, one thickness below 2.9 mm Al and one thickness above 3.5 mm Al to minimize extrapolation in the HVL in calculation) and enter the thickness values in the first column (in mm Al).
6. Insert each thickness between the detector and the treatment head attenuating the entire beam and take an exposure. (adhere the aluminum to the collimator with tape or use a lab stand to hold the aluminum near the head).
7. Record the exposure values in the second column.
Pass/Fail Criteria
The HVL should be above 2.9 mm Al at 80 kVp.
Actions
If the test results is a fail, make every effort to adjust the X-Ray unit to produce a passing result and repeat the test.
If the unit cannot be adjusted to produce a pass, please summarize what actions were taken in the comments section.
Notes
A kVp value must be entered in order for a pass/fail to appear.
Although not prohibited, the use of a survey meter to directly measure HVL is discouraged as this method has been found to be less accurate compared to measurements with aluminum attenuation as described in this procedure. If a survey meter is used to directly measure the HVL, enter the measured value in the comment field provided. You may then ignore the Pass/Fail warning when submitting the results.