## SPM PAN-CBCT
DENTAL SAFETY AND PREVENTIVE MAINTENANCE INSPECTION REPORT
|  |  | (Panoramic / ConeBeam CT) | [10 empty] | Version 1.1 |  |  |  |  | 2021-01-01
EQUIPMENT IDENTIFICATION | [7 empty] | Date of Inspection:
Dental Clinic: | [9 empty] | Room No.:
Address: | [10 empty] | City:
Control Manufacturer: |  |  |  | Model No.: | [7 empty] | Serial No.:
Tube Manufacturer: |  |  |  | Model No.: | [7 empty] | Serial No.:
RADIOGRAPHIC INSPECTION REPORT
A. X-RAY TUBE AND SUPPORT SYSTEM |  |  |  |  | RESULT |  |  |  | REMARKS
Collimator/Filters
Angulation Indicator
Indicator Lights
Security of Mounting (overall)
Stability of Motion (drift)
Electrical Cables/Connections/Grounds
B. CONTROL/TRANSFORMER
Indicator Lights & Meters
Electrical Cables/Connections/Grounds
Overload Protection Circuits
Control Knobs & Buttons
Audible Exposure Indicator
Exposure Control Switch
C. DETECTOR/CASSETTE HOLDER
Beam Alignment
Automatic Exposure Control
Phantom Tests
Image Quality
Ease of Movement
Security of Mounting
D. RADIOGRAPHIC QUALITY | [19 empty] | Result:
kVp/Line Voltage |  | Set Value: |  |  |  | Meas. Value: | [6 empty] | kVp difference < 7%
mA/Line Current |  | Set Value: |  |  |  | Meas. Value: | [6 empty] | mA difference < 10%
Beam Quality (HVL) |  | Set kVp: |  |  |  | Meas. HVL: | [6 empty] | kVp: |  | 70 | 80 | 90 | 100 | 110
[13 empty] | HVLmin: |  | 2.1 | 2.4 | 2.7 | 3 | 3.3
Comments:
E. MANUFACTURER REQUIRED TESTING
Note: Most Panoramic and Cone Beam CT manufacturers prescribe service calibration requirements in their Service/Technical Manuals. Please attach any manufacturer testing results that are not addressed in the SPM form.
Copy of Report Sent to Dental Office?
By submitting this report as a qualified person, I ensure that this ionizing equipment and apparatus is in safe operating
condition, has undergone a radiation calibration, and the results of all testing fall within acceptable tolerances.
SERVICE ENGINEER:
RSU Equipment No.: | [6 empty] | Saskatchewan Journeyman Licence No.:
(to be completed by RSU office)
: //
CLEAR FORM
SAVE SPM FORM
ACCESS

## README
DENTAL SAFETY AND PREVENTIVE MAINTENANCE FORM
Welcome to the RSU dental SPM form. The purpose of the SPM testing is to inform the owner that their X-ray equipment is in safe operating condition. The testing is based on the standards of Health Canada Safety Code 30. Equipment that cannot meet these standards may compromise image quality and/or patient dose. Instructions for each test are provided in the purple tabs of this form. SPM inspections are required under sections 13 and 14 of The Radiation Health and Safety Regulations, 2005. The parts of the legislation that are releveant to dental X-Ray units are outlined below:
Safety and preventive maintenance inspections
13(1) An owner of ionizing radiation equipment and associated apparatus that is
used in a hospital or medical clinic must arrange for the inspection of that
equipment and apparatus by a qualified person in a manner and to a degree that is
satisfactory to an officer to ensure that the equipment and apparatus:
(a) is in safe operating condition; and
(b) has undergone a radiation calibration, the results of which are recorded on a
form supplied by the department.
(2) For the purposes of subsection (1), a person who holds a valid restricted x-ray
journeyperson’s licence issued pursuant to The Electrical Licensing Act is a
qualified person.
(3) An owner of ionizing radiation equipment that is used in a dental, chiropractic
or veterinary clinic must arrange for the inspection of the equipment in accordance
with subsection (1).
(4) A person who conducts an inspection pursuant to subsection (1) or (3) shall,
within 30 days after completing the inspection, submit to the Radiation Safety Unit
on the form mentioned in clause (1)(b) details of all tests carried out and all
measurements made in the course of the inspection.
Frequency of inspections
14(5) Subject to subsections (6) and (7), an inspection required by subsection 13(3) is
to be carried out not less than:
(a) once every three years for dental or chiropractic x-ray equipment; and
(b) once every five years for veterinary x-ray equipment.
(6) No inspection is required pursuant to subsection 13(3) until five years have
elapsed since the date of manufacture of the equipment.
(7) In the case of chiropractic x-ray equipment 15 years of age or older, an
inspection required by subsection 13(3) is to be carried out not less than once per
year.
Complete the form as follows: 1) Fill out the form as completely as possible. - Only the blue fields are required. Calculations are built in. - For each test, either a "PASS" or comment on "FAIL" is required. - Incomplete SPMs will be rejected and returned to the owner. 2) Once complete, hit the SAVE FORM button at the bottom of either page or click "Save As" in the "File" menu. - The save will not complete unless the form is adequately filled out. 3) Submit the saved form via email to radiationsafety@gov.sk.ca. - Submission of the competed form is required within 30 days of inspection.
Track Change Summary
Version | Release Date
1 | 2017-06-01 | u First Release
1.1 | 2021-01-01 | u Updated logo for visual identity
u Changed the date format to mmm dd, yyyy for clarity
u Fixed issue where service engineer info not writing to the database

## Manufacturer Testing
Manufacturer Instructions
Purpose
The purpose of the manufacturer testing is to ensure the panoramic unit is being maintained per the manufacturer's suggestions.
A review of various panormaic/CBCT makes and models has shown that the majority of units have some form of prescribed QA from the manufacturer. Although the testing varies from unit to unit, there are tests that are common to a large percentage of units in use. The common panormaic tests are the following:
u Mechanical Integrity
u Electrical Integrity
u Collimation and Alignment
u kVp Accuracy
u mA Accuracy
In addition, the frequecy recommended by the manufactuer often exceeds the requirements set out in section 14 of The Radiation Health and Safety Regulations, 2005. A subset of the recommended testing requirements of various panormaic and CBCT manufacturers is provided in the table below:
Make | Model | Additional Testing Indicated | Recommended Frequency |  | Source
SOREDEX | Cranex Novus | Yes | Annually | (1) kVp/mA test (2) beam alignment test (3) ball/pin test (4) ground check (5) motor movements (6) position detectors (7) patient position lights (8) mains power supply cable (9) tubehead | Section 7.2 of Users Manual (version 200409) Section 1.5 of the Service Manual (version 200569) Section 4 of the Installation and Set-up Manual (version 200410 rev.2)
| Cranex D |  |  |  | Section 9.2 of Users Manual (version 8201083) Section 9 of the Service Manual (version 8201086)
GE/INSTRUMENTARIUM | OP300 | Yes | Annually | (1) calibration (a) preheat of the tube (b) tube current (c) collimator (d) geometry (e) laser (f) pixel (g) quality check (2) motors and movement mechanisms (3) mains and power supply (4) tubehead (5) ground check (6) fan operation | Section 6.2 of the Installation Manual (version 204389 rev. 6) Section 8 of the Service Manual (version 204904 rev. 3)
BELMONT | X-Caliber EX-1000 | Yes | Semi-Annually | (1) Line Voltage (2) Line Regulation test (3) Tube Current Test (4) Exposure Time Test (5) X-Ray Field Size (6) Alignment Test | Maintenance section of the Operator's Manual
X-Caliber EX-2000
PLANMECA | Promax 3D | Yes | Annually | (1) Peak tube potential (kVp) measurement (2) Half-value layer measurement (3) Exposure time measurement (4) Tube current mA measurement (optional) (5) Feedback signal offset measurement (optional) | Section 2 of the Technical Manual (version 10026634 rev. 6)
SIRONA | ORTHOPHOS XG 3D | Yes | 4, 7, 10, 12, 14… years | (1) Review of records (2) Labels/ serial numbers on the system (3) Functional test (4) Tube data (5) Diaphragm check (6) Cables (7) Safety Related | Inspection and maintenance and safety checks manual (version 6303601)
| ORTHOPHOS XG 3 |  |  |  | Inspection and maintenance and safety checks manual (version 5987685)
| ORTHOPHOS XG 5 |  |  |  | Inspection and maintenance and safety checks manual (version 5987685)
GENDEX | ORTHORALIX 9200 | Yes | Annually | (1) Mechanical Checks (2) Electrical Checks (3) Functional Checks (4) MeasurementChecks (a) verify current (b) verify kV (c) verify dose rate (d) verify reproducibility (e) verify radiation leakage (f) verify earth resistance (g) verify leakage current | Section E of the service manual (version 4519 190 71193)
Procedure
1. Perform the Preventive Maintenance as recommended by the manufacturer.
2. Record the quantitative results along with a Pass/Fail result and the Pass/Fail criteria